## Slide 1: Novel One-Step Centrifugal Sensor System for the Detection of Cyanobacterial Toxin Microcystin-LR
DR. Jenny Fitzgerald – Dublin city university (DCU), Ireland
Aquatic Sciences Meeting
Granada, spain
2015

## Slide 2
The use of recombinant antibodies to detect algal toxins.
Development of a rapid quantitative method for cyanobacterial microcystin
Incorporation onto novel centrifugal immunoassay platform
  - Easy to use
  Low cost
  Portable
  On-site monitoring
Overview

## Slide 3
Algae’s Toxic threat
Agricultural and urban run-off causes increased pollution of lakes, rivers, streams and coastal areas causing eutrophication leading to Harmful Algal Blooms (HABs)
2% of algae produce harmful toxins.
Microcystin is the most ubiquitously occurring cyanobacterial toxin and is present in fresh and brackish waters.
Development of recombinant antibody sensors towards algal-toxins.
Incorporation of antibodies into rapid, ‘easy-to-use’, portable toxin detection device.
Lake Erie, Ohio. August 2014.

## Slide 4
VH
VL
CH1
CL
F(ab')2
VH
VL
CH1
CL
Fab
scFv
IgG
CH1
CL
CH2
VL
VH
VL
VH
VL
VH
VL
VH
VL
VH
VL
VH
Dimeric scFv
Dimeric bifunctional scFv
Development of MC-LR specific Recombinant Antibodies

## Slide 5
Phagemid vector
Resistance gene
ori
Ligation into phage display vector
Transformation in E. coli
Reverse transcription
of mRNA to cDNA
PCR amplification of VH and VL genes
VL
VH
SOE PCR anneals VH and VL
SfiI digest of SOE PCR product
Phage displaying scFv
Recombinant Antibody Library Construction
Extraction of RNA

## Slide 6
0.2
0
0.4
0.6
0.8
1
1.2
Normalised A/A0
MC-LR concentration ng/mL
10
100
1000
10000
5C9
2C5
2C3
2C1
E111
2C4
1F11
2H1
2G1
5A9
Antibodies produced from each round of biopanning – polyclonal ELISA
The most sensitive binder was determined by inhibition ELISA
Recombinant Antibody Library Construction

## Slide 7
The Toxi-sense system Microfluidic System
Notes: Toxi-Sense Box was 3-D printed
Compartment to mount rotor for spinning of disk
Two-channel system to measure test/control zones simultaneously
Fluorescence using low-powered laser (1 milliwatt)
High pass optical filter to measure emission
650nm cut-off and a photodiode
Controlled by a CC2511F32 micro-controller based development board called Wixel.
Fluorescence measured at 1-second intervals
Transmitted to PC via USB

## Slide 8
The Toxi-sense system Microfluidic Disc
Initial three chamber design
-Incubation
-Test
-Control
A loading zone to apply the sample prior to incubation with the antibody was required
A waste zone for collection of flow through
Poly(methyl methacrylate) (PMMA) sheets and Pressure Sensitive Adhesive
Revised five chamber design

## Slide 9
Final Microfluidic platform design
Conical shaped chambers to allow antibodies conjugated on particles to pass through each chamber
Ventilation system covered in PSA to prevent air-flow through the system
Radius of 60mm
Thickness of ~5mm
6 loading zones allow simultaneous detection 6 samples
System lock pinholes to ensure disk immobilisation in correct position
Centre hole of radius 7.5mm for mounting on Toxi-sense platform
This Disc uses manual valves which are aimed to be replaced with automatic valves!

## Slide 10
Microcystin-LR Toxi-Sense
Assay Schematic
ScFv added to device with free antigen
Microcystin
Alexa 647 labelled
Anti-Microcystin ScFv
Control well is coated with anti-chicken IgG

## Slide 11
Toxisense Calibration curve
Chart: Microcystin-LR-OVA concentration (ng/mL)
 | 1366 | 1377 | 1406 | 1414 | 1430 | 1541 | 1870 | 1882 | 2000 | 2047

## Slide 12
Proof of Concept Microcystin-LR detection using Toxisense optical platform
Chart: Free Microcystin (ng/mL)
| 0 | 8 | 40 | 200 | 1000 | 5000
 | 355 | 342 | 335 | 326 | 312 | 303

## Slide 13
Overview of preliminary Toxi-Sense Assay
capabilities
Assay Method | Functional Detection Limits
ELISA | IC50 = 8.1ng/mL LOD = 4ng/mL
SPR | IC50 = 5.9ng/mL LOD = 1.7ng/mL
Fluorescence plate assay | IC50 = 13.7ng/mL LOD = 8.0ng/mL
Fluorescence slide assay | IC50 = 6ng/mL
Toxi-Sense System | LOD = 8 ng/mL

## Slide 14
Conclusion
Results are obtained in less than 15 minutes
Preliminary studies confirm detection of Microcystin-LR under current regulatory limits
Further characterisation and standardisation will confirm assay sensitivity for Microcystin-LR and other variants in complex matrices
Method development underway for on-disk sample processing
Potential for assay multiplexing with other marine toxins
Rapid, low cost, semi-quantitative method for toxin detection
Integrated microfluidic disk
Minimal reagent requirements
Short assay times
Portability-Hallmark of the system
In-situ routine monitoring and on-field screening
High sample throughput-short incubation times
Fast/De-centralisation–low resource areas

## Slide 15
Acknowledgements
Ivan Maguire
Brendan Heery
Dr. Caroline Murphy
Prof. Fiona Regan
Prof. Richard O’ Kennedy
MESTECH
The Applied Biochemistry Group DCU